Light Sheet Manual

Harvard Medical School

Neurobiology Imaging Facility

(NIF)

By Mahmoud el-Rifai  rev.10/2018

Start up:

No Gloves!!

- Turn on the power strip.

Figure 1: power strip

- Turn on the Back-Ups

Figure 2: Back-UPS

- Turn on the computer

- Turn on all laser lines.

Figure 3: Lasers

- Sign in PPMS

- Open “ImSpector Pro” software program.

Note: If laser fails to initialize, then close software and restart software.

PUT ON GLOVES !!!

Check the cleanliness of the objective lens and the Cuvette:

Figure 4: Dipping Cap                                            Figure 5: Cuvette

- Clean the walls of the cuvette with sparkle cleaner or ethanol. (If the walls of the cuvette are not cleaned well the images will be affected).

- Clean the lens of the dipping cap with sparkle cleaner or ethanol with lens tissue not kim wipes.

- Put the air seal ring in the dipping cap, and then place the dipping cap on objective.

Figure 6: Mounting dipping cap.

Mounting the sample:

- Fill the Cuvette with imaging medium. The cuvette has to be filled with the medium up to the recess of the glass

Figure 7:  Never fill the cuvette while it is located in the microscope chamber.

REMOVE GLOVES !!!

- Raise the zoom body all the way up, using the focus knob.

- Rotate the objective out of the working position.

PUT ON GLOVES !!!

- Loosen the hex screw on the right side of the gray stage plate.

- Carefully lift off the gray stage plate.

- Mount sample in the correctly sized sample holder. Try to mount your sample as level and as near the top of the holder as possible.

- Insert the holder into the sample cradle such that the screw of the holder is at a 45° angle to the supports of the cradle. You need the light sheet to be able to pass unobstructed through your sample.

Figure 8: Mounting the sample.

10. Place the cradle into position such that the sample is oriented perpendicular to the light path.  Note: if you will be imaging using a single sheet only, it should be the right side. Therefore, make sure to place your sample such that the most interesting part is facing the right side so that your area of interest is accessible to the sheet: e.g., if interested in spinal cord, place it to the right side so it will not be shadowed in the image.

Note: When removing the sample mount, ensure that the medium does not leak from the cuvette. To safely detach the sample mount, put a tissue under it.

REMOVE YOUR GLOVES!!

- Tighten the hex screw on the right side of the gray stage plate.

- Rotate the objective into the working position.

Finding the sample:

- Set joystick to fast mode by pressing and holding the rightmost button until it beeps.

Figure 9: Joystick

- Select the lowest zoom setting (0.63X) on the microscope and the software.

Figure 10: Magnification setting

- In the software:

- Select the green or red (“OBIS”) laser.

- Set laser power to 10%.

- Set NA to the highest possible setting (optics section).

- Adjust the light sheet width to 20% (optics section).

- Select the right light sheet (light sheet section).

- Click on the “Video” button to begin rapid scanning.

Figure 11: Video button

- Look into the reservoir: you should see the light sheet hitting the sample holder and your sample. If not, maneuver the sample in XYZ until you can see the light sheet going through your sample. Once the light sheet is hitting your sample, lower the Z stage until the light sheet is going through the top of your sample. The light sheet must be in the sample to avoid damage to the system.

- Disengage joystick fast mode by pressing the rightmost button on the joystick once.

- Select the proper laser for your fluorophore.

Figure 12: Laser menu

- Hit the rescale display icon (this is a sun and moon with stars picture on the right side of the imaging window).

Figure 13: Rescale button

- Lower the objective until it is close to touching the imaging medium. Once it is close, lower it more slowly, constantly monitoring the image and adjusting the display to avoid hitting your sample. As you continue to lower the objective, your sample should begin to come into focus.

- Adjust the width of the light sheet to 100% to illuminate the entire field of view evenly.

- Center the sample in the field of view using XY controls. Focus on a structure within the sample using the focus knob on the microscope.

Figure 14: Focus knob

- Go to the top of the sample by rotating the Z knob counter-clockwise.

Note: Never go deeper than 5.2 mm into the sample.

- Set the top of the sample as the end of the Z stack in the xyz-Table Z window.

- Adjust your zoom as desired. Adjust your focus and readjust light sheet width (if needed) after adjusting the zoom. Adjust the zoom setting accordingly in the software.

Autosave settings

Make sure you are writing to the correct directory, with a proper basename. Note that the date and time will be embedded in the name automatically. We highly recommend embedding the zoom you used as part of the basename.

Figure 15: Autosave

Figure 16: Check open series in series …

Imaging setup – the basics

Focusing evenly across the X dimension

If the sample does not seem to be fully in focus across the entire field in the X dimension, you can try the following:

- Use the light sheet from the other side.

- Adjust the horizontal focus position by clicking on the button with two red horizontal triangles to bring up the indicator on the imaging window (measurement tab, Ultra II window). Click and drag the indicator to adjust the position of the horizontal focus until the focus of your structures of interest improves.

Figure 17: Horizontal focus

- Lower the sheet NA. This makes the light sheet dimmer, and thicker in Z in the middle, but more even as a function of X position.

If none of these adjustments leads to an image that looks fully even and in focus across the field of view, try imaging with two light sheets or with multiple horizontal foci.

Exposure and laser power

- Exposure will be the same across multiple channels.

- There are two ways of adjusting laser power:

- Adjust the power of the desired laser using the laser slider in the Ultra II window (move the slider and then hit apply). NOTE: scale is logarithmic and using this mode can lead to slower imaging with more than one channel.

Note: When looking at the sample, it is better to use long exposure (hundreds of ms) and low laser power (10-20%); this avoids bleaching. When imaging the sample, it is better to use very short exposures (10 ms) and high laser power; this speeds up the acquisition.

Imaging with multiple horizontal X sheet foci:

You can set the start and stop position for the horizontal focus and choose how many images you

would like to capture as the focus moves across the field.

- Check that the zoom factor in the software is the same as on the microscope knob.

- If this is not set properly, the dynamic focus will give very poor results.

- Set dynamic focus to “ON” (horizontal focus section, measurement tab, Ultra II window).

- Click on the icon with multiple horizontal triangles and set the area in the image over which you want to move the focus.

- Input the recommended number of focus positions.

- Enter the desired exposure time; ensuring that it is higher than 100 ms (the software will adjust it slightly automatically).

Cropping

- If you only care about a subregion in the field of view, cropping will significantly reduce your file sizes and save you transfer and processing times later on.

- To crop, draw an ROI on the image and use the “Box Inside a Box” button. To ensure optimal quality, your ROI should be centered in the field of view.

- To uncrop, click on the “Full Field” button in the Ultra II window (Measurement tab, sample section)

Acquiring Z stack – 1 channel:

- Check that the zoom factor in the software is the same as on the microscope knob. • If this is not updated properly before you begin imaging, your metadata will be incorrect later.

Figure 18

- While imaging in video mode, move to the top of your sample (move the joystick counterclockwise).

In the Table Z dialogue box:

- Click the “Set as Zero” button.

- Click the End target button.

- Move to the bottom of your sample (move the joystick clockwise).

DO NOT EXCEED 5.2 mm deep in the sample with the corrected dipping cap.

Figure 19

- Click the Start target button.

- Set Z step size. 1/2 or 1/3 of the light sheet thickness is a good starting point, but the ideal step size will depend on sheet NA and the size, spacing and desired imaging detail of the features of interest in the sample. (5um for .63, 1um for 6.3)

- Stop Video Mode.

- Add “xyz-Table Z” to the device list in the measurement wizard window.

Figure 20

- Make sure that “autosave” is on.

- Press START. Measurement Wizard displaying designated Z-stack properties before clicking “Start” for image acquisition.

Figure 21

Acquiring Z stack – more than 1 channel:

In addition to steps 1-9 above:

- Check the boxes for the filter sets you would like to use.

Figure 2

The chromatic correction button  next to the filter box.

Figure 23

- Go to your other channels and check focus in each one. Do not move the focus knob. If the focus is off, use the chromatic correction slider to adjust focus. Once it is properly positioned, click the “Set” button to save the position.

- Note that higher magnifications require more movement of the chromatic correction lens. Also note that some combinations of fluorophores and magnifications cannot be corrected (for example, 647 and 488 at 6.3X zoom).

- Adjust the laser power for each channel.

- You can adjust the power of each laser using the laser slider in the Ultra II window. Choose a channel by highlighting it (not checking it) and make the proper adjustment, then click “Apply” for each. Note that if you use different laser powers for different channels, imaging will be slower, particularly if channels are acquired sequentially at each Z plane.

- Adjust the exposure for the camera. Note that there is only one exposure setting for all the channels.

- Add “Ultra Filter” to the device list in the measurement wizard window.

Figure 24

- Note that the order in the device list is the order in which things will be done. For example, with two channels, Ultra II Filter 1st, Table Z 2nd would go to a Z position, take images in each channel, and then move to the next Z position; the opposite order (Table Z 1st, Ultra II Filter 2nd) would take a complete Z stack in the first channel, then switch to the second channel and take a complete Z stack for that one.

- Make sure that “autosave” is on for all devices.

- Make sure that the Ultra Filter setting has the “split” engaged.

- Press START.

Generating a Mosaic Image:

Notes:

• Be careful with the objective bumping into the edge of the travel range when tiling at high magnifications.

In addition to steps 1-9 (Z stacks), and 1-7 (multiple channels) above:

- Choose measurement mode multi-color mosaic acquisition.

- In xyz-Table Visual XY, select the “Mosaic” tab.

Figure 25

- In set parameters, under advanced options, set the overlap between 10%-20%.

Note: Overlap depends on your magnification and your sheet width, higher magnification and a thinner sheet width requires more overlap up to 40% sometimes.

- While imaging in video mode, adjust the position of the stage to one corner of where you want mosaic.

- This can be done using the joystick or by dragging the red box in the XY Table window.

- Double click the red box, then click and drag it to create area for your mosaic.

- The size of the mosaic area can be adjusted by dragging the sides of the box delineating the mosaic area.

- You can also increase the size of the mosaic by dragging the red box to an area outside of its borders and double clicking.

- To generate an overview of the mosaic area:

- Remove all devices that will not be included in the overview image from the device list (Z certainly, possibly Ultra Filter).

- Add “xyz-Table X” and “xyz-Table Y,” in this order, to the device list in the measurement wizard window.

- Deactivate autosave for each device in the list by clicking the small disk in the “AS” column next to it. The disk should now be greyed out.

- Activate the split view for xyz-Table X and xyz-Table Y by checking the box next to them (at the top of this column, “split” is written).

- Hit the “Start” button to begin imaging.

- After imaging, adjust the display in one tile and hit CTRL-Q to equalize those settings to all tiles.

- If the mosaic looks good, proceed with complete imaging.

- Decide on the imaging order. A typical example would be to do Z first, then channels, then mosaic positions. The order of the devices in the device list should reflect this order. Note that for the tiling, you will need to add xyz-Table X and xyz-Table Y, in this order.

- When ready for final acquisition, check that all channels have autosave and that—if used—Ultra Filter has the “split” option engaged.

- Press START.

In the name of the image, the first parenthesis is the Y value and the second is the X value [YY x XX], even though that is the reverse order of that used when adding it to the device list.

Viewing data in ImSpector (crashes, prefer Imaris or Fiji):

- Open “Series Viewer.” This is the icon along the top that has three squares, two arrows and a line on it.

- In the window that opens, select the folder containing the dataset you wish to view.

- Select the dataset in the “Image Series Found” part of the window.

- Click the down arrow to add that file to the bottom box of the window.

- Click the “Load” button.

- Use the “Series Viewer” window to move through your images, not the bars along the image.

- Close the image window when you are done.

Changing samples:

- Raise the zoom body all the way up, using the focus knob.

- Rotate the objective out of the working position

PUT ON GLOVES!!!

- Remove sample cradle and move it to the hood. Use paper towels to avoid imaging medium spills.

- If the dipping cap will be out of the medium for more than couple minutes, you should dip the cap in the medium while mounting your new sample. Otherwise, the dipping cap will dry then you have to wash it again.

- Swap sample.

- Insert the holder into the sample cradle such that the screw of the holder is at a 45° angle to the supports of the cradle.

- Place the cradle into position such that the sample is oriented perpendicular to the light path.

REMOVE YOUR GLOVES!!

- Rotate the objective into the working position.

Shut down:

- Raise the zoom body all the way up, using the focus knob.

- Loosen the hex screw on the right side of the gray stage plate.

- Rotate the objective out of the working position.

PUT ON GLOVES!!

- Wipe off all imaging medium from the objective using lens paper.

- Remove the cradle that holds the sample with a paper towel.

- Carefully lift off the gray stage plate. Put the reservoir’s black lid back on.

- Moisten a paper towel with ethanol and clean off any imaging medium from the stage or gray plate. Discard the towel contaminated with imaging medium in hazards waste container.

- Put away your sample and clean the sample holder, cradle and tray in the sink with water.

- Remove the objective lens; gently wipe the lens with sparkle glass cleaner with lens tissue.

Figure 26

REMOVE YOUR GLOVES!!

- Exit the software.

- Log off computer.

- Turn off all lasers.

- Turn off the power strip.